This checklist is ONLY intended for use with terminally sterilized devices in their final packaging applying H2O2 sterilization. Please use the alternative checklists for H2O2 sterilized re-processable devices (MED_T_09.82) or generalized / non-standard sterilization methods (MED_T_09.80). Please request a current copy from your TÜV SÜD Product Service primary contact.

Application ID (as it appears in the application form / change notification form)

- [X] in this document indicates a document to be named including page number – submitted for evidence. Grey text (for guidance) may be replaced/deleted.

- In case of a Change Notification, please only fill in the applicable sections.

# Short Product Description relevant for H2O2 Gas Plasma Sterilization

Note: Please replace grey italic text with respective information

Short description incl. picture of the device - in case of changes, as far as relevant

Description of the device as far as relevant for sterilization (pictures for clearer understanding): To be added Product Schematic and / or Photo of product, size, material, Intended Use / Intended Purpose according to IFU (inclusive total application duration, body contact, implantable, patient group), Packaging description, Picture Variants under assessment: To be added Product variants (e.g. Same product in different SBS, Multiple products in same SBS) Description of the Sterile Barrier System Specifications used at sterilization - What is the representative / worst-case SBS for the product under assessment?

| Manufacturing facility and certification status of the applicable sterilization sites / facilities |  |
| --- | --- |
| (Product) Manufacturing site(s) to be named | Please provide the applicable QMS Certificate 13485 of all applicable manufacturing site(s), if not already listed on the certificate. |
| Sterilization site / Validation contractor | Please provide the applicable QMS Certificate 13485 of the used sterilization site In case of identical laboratory subcontractor (e.g. residual analysis, microbial testing) please additionally provide ISO 17025 certificate including respective scope. In case the sterilization site and the H2O2 sterilization study provider are not identical, please provide respective information for all involved parties. |

| External laboratories if used for sterilization validation and certification status of the laboratory |  |
| --- | --- |
| Name of the laboratory | Please name the test done by the laboratory (e.g. Microbiology BI testing, Sterility testing, Bioburden, Residuals). Please provide the applicable QMS Accreditation Certificate (e.g. ISO 17025 or GLP) |
| Name of the laboratory | Please name the test done by the laboratory (e.g. Microbiology BI testing, Sterility testing, Bioburden, Residuals). Please provide the applicable QMS Accreditation Certificate (e.g. ISO 17025 or GLP) |

# Production related Information

## Used Hardware Specifications

Note: Please replace grey italic text with respective information. Please add additional lines if required.

| Equipment including Identifier (e.g. int. ID/ serial number) | Site / Location | Applicable cycle operated by the equipment for the device in question | Typical load configuration | Usable chamber volume in m3 | All sensors / measurement devices (internal + external sensors, dataloggers for validation) are calibrated |
| --- | --- | --- | --- | --- | --- |
| e.g. Super H2O2 Plasma | Inhouse or external source | Please name the cycle and Version/ Revision of cycle e.g. FAST & CLEAN please list ALL cycles that are used together with the product | e.g. Flexible Endoscope / Central Sterilization mixed load | 5,4 | ☐ Yes ☐ No |

Notes:

- Sensor calibration and equipment maintenance need to cover all critical sensors, e.g. optical / spectral detector for H2O2 as in-process monitor.

- Sterilizer equipment used for the validation must be CE certified by a European Notified Body. In any other case a full sterilizer equipment design documentation needs to be submitted to be reviewed additionally.

| Sterilizing agent definition |  |
| --- | --- |
| Clear identifier, e.g. H2O2 prefilled cartridge | Further details on the sterilizing agent presentation: Half-Dose cartridge, Booster, 3rd party cartridge, fill volume, sample picture, REF Code or order code from supplier, other unique identifier, certificate of conformity or certificate of analysis In case during validation the H2O2 is manually injected (e.g. by syringe) into the process chamber, please provide above data for the separately sourced H2O2 bottle and how the H2O2 is considered to be equal / comparable to the routine H2O2 in cartridges. Above information may be used as guidance and not each individual point may be available for the individual submission. |

Accessories & Components as part of the validation and routine processing

Please list all accessories and components that are to be used in preparation and/or inside the sterilization chamber for the validation activities and also for intended routine sterilization. - ExtraDeep Instrument Tray - CompatibleTRAY Instrument Tray - SuperFoil Sterilization Wrap - Ultra PRECISE Chemical indicator Tape

Biological Indicator (BI) / Process Challenging Device (PCD) description

Minimum expected level of information: - Detailed description of the BI (carrier material, dimensions, microorganism, population, D-Value (H2O2), survival/kill window, dose/response characteristics, adherence to EN ISO 11138-1, Certificate of Analysis, Certificate of Compliance) In case it is deviated from the standard microorganism Geobacillus stearothermophilus, a full investigation study (including scientific data) needs to be provided to demonstrate equal or higher resistance. If biological indicator is not sourced from a certified supplier (e.g. manufactured by internal microbiology lab), then a full design documentation of the BI manufacturing process needs to be added to the documentation or separately covered by a supplier audit of this manufacturing process. - Detailed description of the PCD (picture / drawing, placement of BI inside PCD, inoculation with spore suspension) - Is the used ePCD available in the market to the commercial end-user? If not, how is the correlation (e.g. D-Value, resistance characteristics) between the publicly available ePCD (for routine release) to the custom ePCD within this validation achieved?

## If Applicable: Cleaning of Product in Manufacturing before Sterilization

Note: Please replace grey italic text with respective information. Please add additional lines if required.

| Cleaning process description |  |
| --- | --- |
| Are cleaning parameters that are specific for the medical device defined? | ☐ yes documented in [X,p.y] ☐ no please justify: Please describe the cleaning process: Parameters of the cleaning process may be specific load configuration, positioning, connection, accessories, process chemicals, pressures or temperature limit(s)… |
| Have cleaning studies been performed in validation of the used equipment to approve the respective cleaning process step is able to deliver appropriate performance? | ☐ yes documented in [X,p.y] ☐ no please justify: Cleaning studies shall reflect the potential contamination and evidence of reproducible elimination of the contamination to a safe level (see EN ISO 17664) |
| Are (cleaning) process residuals within limits? | ☐ yes documented in [X,p.y] ☐ no please justify: Residuals like endotoxins, particles, org. inorganic contaminations, detergent residues with adequate risk related to the device and body contact shall be safe and acceptable. |
| Are preventive maintenance operations defined including frequencies? | ☐ yes documented in [X,p.y] ☐ no please justify: Examples for preventive maintenance are: exchange of cleaning media and/or equipment |
| Residual water / cleaning agent impact on H2O2 gas plasma sterilization in vacuum | Please specify maximum allowable limits for residual water / residual cleaning agent on the device-to-be-sterilized that could adversely affect the subsequent sterilization cycle. Only applicable if no disinfection step follows this cleaning step. Otherwise refer to the same question in the next section 2.3 |

## If Applicable - Disinfection of Product in Manufacturing before Sterilization

Note: Please replace grey italic text with respective information. Please add additional lines if required.

| Disinfection process description |  |
| --- | --- |
| Are parameters that are specific for the medical device defined, such as specific load configuration, positioning, connection, accessories, process chemicals, pressures or temperature limit(s)? | ☐ yes documented in [X,p.y] ☐ no please justify: Please describe the disinfection process: Parameters of the disinfection process may be specific load configuration, positioning, connection, accessories, process chemicals, pressures or temperature limit(s)… |
| Were disinfection studies performed in validation of the used equipment to approve the respective disinfection process step is able to deliver appropriate performance? | ☐ yes documented in [X,p.y] ☐ no please justify: Disinfection studies shall reflect the potential contamination and evidence of reproducible elimination of the contamination to a safe level (see EN ISO 17664) |
| Are process residuals within limits? | ☐ yes documented in [X,p.y] ☐ no please justify: Residuals like endotoxins, particles, org. inorganic contaminations, detergent residues with adequate risk related to the device and body contact shall be safe and acceptable. |
| Are preventive maintenance operations defined including frequencies? | ☐ yes documented in [X,p.y] ☐ no please justify: Examples for preventive maintenance are: exchange of disinfection media and or equipment |
| Residual water / cleaning & disinfection agent impact on H2O2 gas plasma sterilization in vacuum | Please specify maximum allowable limits for residual water / residual cleaning & disinfection agent on the device-to-be-sterilized that could adversely affect the subsequent sterilization cycle. |

## Only applicable to Change of Clean Room Control / Validation

Note: Please replace grey italic text with respective information. Please add additional lines if required.

| Cleanroom | Please identify the cleanroom(s) where the manufacturing takes place, including ISO classification |
| --- | --- |
| Are action and alert levels/limits set appropriately for the subsequent product bioburden in cleanroom processes? | ☐ yes documented in [X,p.y] ☐ no please justify: Acceptance criteria for “in operation” condition: Airborne particles [size]: particles/m3 Airborne microbiological contamination: cfu/m3 (and/or settle plates) Surface microbiological contamination: cfu/ surface area Product bioburden: cfu (type – spores, fungi, anaerobe, bacteria) The bioburden shall be known to a degree to make decisions on resistance |
| Monitoring points are defined for the above-mentioned measurements | ☐ yes documented in [X,p.y] ☐ no please justify |
| Was IQ, OQ, PQ of the cleanroom successfully established? | ☐ yes documented in [X,p.y] ☐ no please justify: |
| Is all measuring equipment in a calibrated state? | ☐ yes documented in [X,p.y] ☐ no please justify: |
| Are utilities and media under surveillance | ☐ yes documented in [X,p.y] ☐ no please justify: Please specify what media and related acceptance criteria are defined. e.g. for water, compressed air… |
| Are environmental parameters defined: | ☐ yes documented in [X,p.y] ☐ no please justify: Please specify - where applicable: Temperature: Humidity: Pressure gradient/Pressure level at each room: Air exchange rates: Recovery time: |

## Cycle Specification

Note: Please replace grey italic text with respective information for inhouse and outsourced processes. Please add additional lines if required.

Please paste copy of the cycle specification used at routine sterilization:

Please paste or reference [X] cycle specification here In case of using a different set of parameters for the process validation, please explain the differences in detail and provide a rationale for the deviation (see also section 2.8. for validation cycle specifications). In case only pre-defined sterilization programs / cycles can be selected by the end-user, please provide a rationale and reference to the cycle descriptions in the equipment manufacturer’s manual.

| Is re-sterilization allowed? | ☐ Yes Please state the maximum number of allowed re-sterilizations of the product or product family under assessment. How was this maximum number of possible re-sterilization cycles determined? Is a study available that supports this value? ☐ No, please justify |
| --- | --- |
| Was product functionality and residual limits verified after maximum amount of allowed sterilization cycles | Besides product functionality testing please also consider the possible amount of H2O2 (or other) process residues after maximum amount of re-sterilization cycles. ☐ Yes, documented in [x] ☐ No, please justify: |

## Basic Validation Development Data

Note: Please replace grey italic text with respective information.

| Validation Method | ☐ Overkill half cycle / partial cycle approach (EN ISO 14937 Annex D / Approach 3) ☐ BI and bioburden (EN ISO 14937 Annex C / Approach 2) ☐ Bioburden (EN ISO14937 Annex B / Approach 1) ☐ Product adoption to an existing cycle: Please assure that the adoption rationale and MPQ/PPQ validation data are submitted for the predicate device. This data is documented in [X] ☐ Other: Please specify and add rationale for not using a standardized validation approach |
| --- | --- |
| Validation related Procedures | Procedures to be included as part of the submission: ☐ How new product will be added to a sterilization cycle- is documented in [X] ☐ By what events a new validation is triggered and what step of validation has to address what type of measurements - is documented in [X] Please provide the review interval of data, and interval of time till repeat MPQ PPQ studies. When was the last MPQ, PPQ study executed? What are further criteria that trigger a revalidation study (e.g. Product changes…)? Is a routine interval specified (e.g. annual requalification)? ☐ How new equipment/process will be qualified/validated - is documented in [X] ☐ What significant changes related to sterilization will be notified to the Notified Body - is documented in [X] ☐ Validation protocol for the actual sterilization cycle including acceptance criteria - is documented in [X] |

## PPQ Physical Performance Qualification

Note: Please replace grey italic text with respective information for inhouse and outsourced processes. Please add additional lines if required.

| Please specify the validation load configuration: | Please specify the load used during validation at PPQ and MPQ: Please consider min and max configuration in case of widely varying load configurations, scheme of total load), number of BIs, number of Sensors , scheme of position of BIs, total load volume, density, amount of adsorptive material. The data is documented in [X] |
| --- | --- |
| Please specify which product was used in the load: | The product is the same as in section 1 ☐ yes ☐ no – Please provide a description and justification |

Note:

- Please include additional information on how the used load / dummy load is representative for the expected routine load configuration

## MPQ

Note: Please replace grey italic text with respective information for inhouse and outsourced processes. Please add additional lines if required.

Please paste copy of the cycle specification used at validation of sterilization MPQ (e.g. half-cycle / partial cycle):

Please paste or reference [X] here. Cycle record summaries are documented in [X] If identical to routine cycle as described / referenced in section 2.5, please specify.

| MPQ Processing - Microbial Performance Qualification |  |
| --- | --- |
| Worst-case sterilizer conditions met? | Please explain, how the worst-case sterilizer process limits have been challenged (e.g. using H2O2 cartridge at end-of-shelf-life, highest tolerable process pressure, lowest plasma power). Please only consider parameters, that are changeable by the equipment end-user. If pre-defined cycle programs & parameters are unchangeable by the end-user only maximum possible load configuration is considered relevant. This ensures that the sterilizer equipment, operating at the lower process limits, is still capable of achieving an SAL<10-6 in the chosen cycle for the device-under-test. |
| Placement scheme of BI and/or PCD within the load and in/on the actual device | Please provide detailed information where the BI / spore inoculation / PCD placement in or on the medical device is done, including a rationale (e.g. longest lumen, difficult to reach crevice, high absorbent material, high surface roughness, materials that catalytically dissociate H2O2 on the surface, etc.). Additionally, please provide the placement of BI/PCD within the chamber load to resemble routine conditions (load release). |
| BI incubation conditions (Reduced Incubation Time RIT) | If incubation duration is less than 7 days, please provide a reduced-incubation-time (RIT) study in a representative load (no manufacturer study in a BIER vessel) to substantiate the incubation time [X]. |
| BI results including culture conditions and time between end of cycle and BI testing | ☐ BI results including culture conditions and time between end of cycle and BI testing is documented in [X] |
| Confirm resistance hierarchy from least resistant (natural product bioburden) < IPCD <= ePCD (most resistant) | ☐ sublethal / reduced cycle for BI resistance was performed and is documented in [X] ☐ Data on the relationship of the resistance between ePCD, iPCD, worst-case Products, natural Bioburden is provided and is documented in [X]. |
| Material-dependent microbial resistance | Please state how material-dependent microbial resistance is covered during the validation activity? It is well-known, that different material surfaces exhibit large variability in microbial inactivation time. Please consider the following - What would be the worst-case material with respect to the sterilizing agent? - Is this material covered / included by the validation activity? - How is the material covered (e.g. direct inoculation on the actual device, BI using this material as carrier material, representative worst-case inoculated dummy material, etc.) - Has this been pre-established by the equipment manufacturer in an internal study that could be referenced / audited? |
| Pre-sterilization Bioburden | Please specify the bioburden level and acceptance limits e.g. with respect to bacteria, yeast/molds and anaerobic bacteria (were these investigations at least part of the initial assessment of bioburden?). Please provide the bioburden trending data as a summary of the past two (2) years, if not available at least for the validation LOT plus a rationale (e.g. manufacturing occurs only intermittently = twice per year). |
| Is endotoxin testing applicable for the device under assessment? | ☐ no ☐ yes [X,p.y] Please specify the method and related results. The data is documented in [X]. (e.g. in case of direct contact to blood, CNS, eye or other systemic exposure) If endotoxin monitoring is not required / not executed, please provide a detailed rationale. |

Please assure that the following phases and process values and tolerances are part of the overall validation requirement.

Please be aware that the below parameter-list is not exhaustive to cover the cycle and load types but are often omitted causing deficiencies and are therefore specifically requested.

| Phase | Acceptance criteria: values and tolerance | Results measured and within tolerances EN ISO 14937:2009 item 9.4.4 | Comments if needed |
| --- | --- | --- | --- |
| Total Cycle time | Min / Max running time | (best provided in a table showing setpoint/tolerances against measured data) |  |
| Chamber wall temperature | Min / Max. temperature | See above |  |
| Diffusion time | Min / Max diffusion time of the sterilizing agent | See above |  |
| Plasma time | Min / Max. plasma running time | See above |  |
| Plasma pressure | Min / Max. allowed pressure during the plasma phase | See above |  |
| Plasma power | Min / Max. allowed plasma power | See above |  |
| No. of injections per cycle | Exact number of H2O2 injections into the chamber | See above |  |
| Estimated H2O2 concentration within the sterilizer chamber at dwell phase | ~ xxx mg/liter | See above |  |
| Total H2O2 Injection volume for all injections | ~ 5 x 10g / ml | See above |  |

Note:

- Please provide above table for each cycle that is to be considered. In this case, please copy above table and fill out the fields for each individual cycle.

- A reference to the detailed cycle specification in the documentation is also possible.

# H2O2 Residuals   (EN ISO 10993-17)

Note: Please replace grey italic text with respective information for inhouse and outsourced processes. Please add additional lines if required.

|  |  |  | Comments |
| --- | --- | --- | --- |
| Is a rationale provided for selection of representative sample(s) | ☐ Yes, documented in [X] | ☐ No, please justify: | To be added if any |
| Are the applicable Allowable limits for H2O2 residuals specified (including max. re-sterilization cycles) and confirmed for the product in question considering the patient population provided e.g. by TCL, TI, TTC, | ☐ Yes, documented in [x] | ☐ No, please justify: | To be added if any |
| Results and Test method provided Method qualification can be supported by ISO 17025 certificate of an accredited laboratory. | ☐ Yes, documented in [x] | ☐ No, please justify: | To be added if any |

# Routine Processing

Note: Please replace grey italic text with respective information

| Routine Processing |  |
| --- | --- |
| Routine release | ☐ BI and Physical load parameters Placement scheme and number of BIs, Certificate of BI or spore suspension ☐ BI incubation time 7 days / 14 days? If incubation duration is less than 7 days, please provide a reduced-incubation-time study in a representative load (no manufacturer study in BIER vessel) to substantiate the incubation time [X]. ☐ Parametric release strategy provided in document [X] Please provide a detailed parametric release investigation in order to demonstrate process robustness within the key process parameters, including the clear definition of tolerances / tolerance bands for each of the key process parameter. The study is presented in document [X] |
| Placement scheme and number of BIs, BI certificate provided? | ☐ Yes, documented in [x] ☐ No, please justify: |
| Load configuration | ☐ Dedicated load Product configuration is fixed in number and location within chamber ☐ Mixed load different products allowed Please add information on min/max load variation, if applicable: |

| Release by client: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Date |  | Signature |  | Full Name |
|  |  |  |  | Name of Legal Manufacturer |

Note as to the signature’s relevance: If this document is officially signed, the provided rationales and data herein can be officially used by the reviewer. Otherwise, only the referenced documents can be used as evidence.